|  |
| --- |
|  |
| **New formulation for AST and ALT tests****Notification Date: 8/27/2025****Effective Date: 9/9/2025** |

**SUBJECT:***Effective 9/9/2025 – St. Mary’s Madison Hospital will begin using a new formulation for AST and ALT liver function tests.*

In-house validation studies demonstrate a positive bias compared to the previous method. Additionally, AST has a new adult reference range of 0-48 U/L (previously 5-34 U/L).

Please interpret changes in values from the old method with caution and re-baseline patients as necessary.

Direct questions or comments to:
    Molly Gurney, M.D.
    Assistant Laboratory Medical Director, Chemistry
    SSM Health St. Mary’s Hospital Madison
    Office: (608) 258-6914
    molly.gurney@ssmhealth.com

Thank you,
***Adam Morgan, M.D.
Pathology, SSM Health St. Mary’s Hospital Madison****700 South Park Street
Madison, WI 53715
Office: (608) 258-6914*